## Mode Change
Enter to | StandbyMode | BreathingMode | NonVenMode | SelftTestMode | MEMode | NormalBreathingMode | AlarmBreathingMode | OcclMode | NoGasMode | HFOMode | N_CPAPMode | SponMode | MixMode | AcMode | PCV_CPAPMode | APRVMode | ApneaBackupMode | NIV_Mode | PCV_ACMode | PCV_SIMVMode | VCV_AC_Mode | VCV_SIMV_mode | VA_AC_mode | ApneaVAMode | VA_SIMV_mode | AvMode | DisconnectionMode | HardwareErrorMode | HFOExtremeLowMapMode | HFOHighMap_2Mode | HighPressMode
StandbyMode | x | NA | x | x | x | NA | NA | x | x | NE | x | phaseMgrPtr->ProcessNewPhase (eExhPhase) | phaseMgrPtr->ProcessNewPhase (eMandInhPhase) triggerMgrPtr->EnableTrigger (eBpTimeTrigger) | phaseMgrPtr->ProcessNewPhase (eMandInhPhase) triggerMgrPtr->EnableTrigger (eBpTimeTrigger) | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
BreathingMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
NonVenMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
SelftTestMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
MEMode |  | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
NormalBreathingMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
AlarmBreathingMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
OcclMode | x | NA | x | x | x | NA | NA | x | x | NE | x | phaseMgrPtr->ProcessNewPhase (eExhPhase); | phaseMgrPtr->ProcessNewPhase (eMandInhPhase) triggerMgrPtr->EnableTrigger (eBpTimeTrigger) | phaseMgrPtr->ProcessNewPhase (eMandInhPhase) triggerMgrPtr->EnableTrigger (eBpTimeTrigger) | x | x | x | x | x | x | x | x | x | x | x | phaseId = phaseMgrPtr->GetCurrentPhaseId () The mode change should only occur in exhalation Deliver a mandatory inhalation Enable the breath period trigger | x | x | NE | NE | x
NoGasMode | x | NA | x | x | x | NA | NA | x | x | NE | x | phaseMgrPtr->ProcessNewPhase (eExhPhase); | phaseMgrPtr->ProcessNewPhase (eMandInhPhase) triggerMgrPtr->EnableTrigger (eBpTimeTrigger) | phaseMgrPtr->ProcessNewPhase (eMandInhPhase) triggerMgrPtr->EnableTrigger (eBpTimeTrigger) | x | x | x | x | x | x | x | x | x | x | x | phaseId = phaseMgrPtr->GetCurrentPhaseId () The mode change should only occur in exhalation Deliver a mandatory inhalation Enable the breath period trigger | x | x | NE | NE | x
HFOMode |  | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
N_CPAPMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
SponMode | x | NA | x | x | x | NA | NA | x | x | NE | x | ASSERTION(phaseMgrPtr->GetCurrentPhaseId() == eExhPhase); | The mode change should only occur in exhalation. Determine the time elasped since the start of the last inh phase breathPeriod = triggerMgrPtr->GetTimePeriod (eBpTimeTrigger) When entering the Mixed Mode from any breathing mode, determine the time elapsed since last inhalation. | The mode change should only occur in exhalation. Determine the time elasped since the start of the last inh phase The value associated with the Breath Period Time Trigger is the breath period. Retrieve it. When entering A/C from any other mode, figure out time elapsed since the start of the last inhalation. | x | x | x | x | x | x | x | x | x | x | x | phaseId = phaseMgrPtr->GetCurrentPhaseId () The mode change should only occur in exhalation Deliver a mandatory inhalation Enable the breath period trigger | x | x | NE | NE | x
MixMode | x | NA | x | x | x | NA | NA | x | x | NE | x | ASSERTION(phaseMgrPtr->GetCurrentPhaseId() == eExhPhase); | The mode change should only occur in exhalation. Determine the time elasped since the start of the last inh phase breathPeriod = triggerMgrPtr->GetTimePeriod (eBpTimeTrigger) When entering the Mixed Mode from any breathing mode, determine the time elapsed since last inhalation. | The mode change should only occur in exhalation. Determine the time elasped since the start of the last inh phase The value associated with the Breath Period Time Trigger is the breath period. Retrieve it. When entering A/C from any other mode, figure out time elapsed since the start of the last inhalation. | x | x | x | x | x | x | x | x | x | x | x | phaseId = phaseMgrPtr->GetCurrentPhaseId () The mode change should only occur in exhalation Deliver a mandatory inhalation Enable the breath period trigger | x | x | NE | NE | x
AcMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
PCV_CPAPMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
APRVMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
ApneaBackupMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
NIV_Mode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
PCV_ACMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
PCV_SIMVMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
VCV_AC_Mode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
VCV_SIMV_mode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
VA_AC_mode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
ApneaVAMode | x | NA | x | x | x | NA | NA | x | x | NE | x | ASSERTION(phaseMgrPtr->GetCurrentPhaseId() == eExhPhase); | The mode change should only occur in exhalation. Determine the time elasped since the start of the last inh phase breathPeriod = triggerMgrPtr->GetTimePeriod (eBpTimeTrigger) When entering the Mixed Mode from any breathing mode, determine the time elapsed since last inhalation. | The mode change should only occur in exhalation. Determine the time elasped since the start of the last inh phase The value associated with the Breath Period Time Trigger is the breath period. Retrieve it. When entering A/C from any other mode, figure out time elapsed since the start of the last inhalation. | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
VA_SIMV_mode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
AvMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
DisconnectionMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
HardwareErrorMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | ModeMgr::S_GetInstance()->GoToOperatorSetMode(); | x | x | x | x | x | x | ModeMgr::S_GetInstance()->GoToOperatorSetMode() | x |  | x | x | NE | NE | x
HFOExtremeLowMapMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
HFOHighMap_2Mode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | x | x | x | x | x | x | x | x | x |  | x | x | NE | NE | x
HighPressMode | x | NA | x | x | x | NA | NA | x | x | NE | x |  |  |  | x | x | ModeMgr::S_GetInstance()->GoToOperatorSetMode(); | x | x | x | x | x | x | ModeMgr::S_GetInstance()->GoToOperatorSetMode() | x |  | x | x | NE | NE | x
Default | Close PurgeFlow Close OilPump Disable HFO InpedanceLine->SetState(eOn) InpedanceLine->SetDesiredState(eOff) SafetyValve->SetDesiredState(eOpen) Disable all trigger Enable mode change trigger PhaseMgr::S_GetInstance()->ProcessNewPhase(eDiagnosticPhase) PInhAzTestPtr->SuspendTest(); PExhAzTestPtr->SuspendTest(); blank all patient data when enter standbymode HfoValve->DisableMotorErrorCheck() O2Valve->DisableMotorErrorCheck() AirValve->DisableMotorErrorCheck() ExhValve->DisableMotorErrorCheck() WatchDog::S_GetInstance()->DisableWatchDog() | NA | Open the safety valve Initialize pointer to the TriggerMgr Put the ventilator into the alarm phase Disable all triggers triggerMgrPtr->EnableTrigger (eUpdateStatusTimeTrigger, INIT_UPDATE_STAT_TIME) Suspend auto-zero scheduling | Put the ventilator into the SelfTest phase Disable all triggers InhPressureZero->SetDesiredState(eOff) ExhPressureZero->SetDesiredState(eOff) SafetyValve->SetDesiredState(eOpen) Suspend auto-zero scheduling | Enable the mode change trigger Process DiagnosticPhase Suspend auto-zero scheduling. | NA | NA | AlarmMode::Enter (fromModeId) Enable the reset occlusion trigger Update the Occlusion alarm status | AlarmMode:Enter(fromModeId) Enable the reset gas loss trigger Enable the gas loss trigger If there is a low O2 supply pressure condition, set the total gas loss alarm to the active state If the blower has never been stopped, it indicates that this mode is being entered due to the loss of the air source for the first time | NE | enable trigger PhaseMgr::S_GetInstance()->ProcessNewPhase(eNCPAPPhase) NCPAPController->ResetIMVStaticVariables() AirFlowController->Enable() O2FlowController->Enable() NCPAPController->Enable() SafetyValve->SetDesiredState(eClose) | UpdateCommonTriggers () ReportModeChangeInducedReset() (Hip2AlarmTrigger::S_GetInstance ())->OverrideTrigger (eTrue) | UpdateCommonTriggers () triggerMgrPtr->EnableTrigger (eExhHoldOpReqTrigger) Set the window to mandatory | UpdateCommonTriggers () triggerMgrPtr->EnableTrigger (eExhHoldOpReqTrigger) ReportModeChangeInducedReset () | if(fromModeId != eApneaBackupMode) SpontMode::Enter(fromModeId) PhaseMgr::S_GetInstance()->ProcessNewPhase(eExhPhase) | (PhaseMgr::S_GetInstance())->ProcessNewVentType(ePressureVentilation) CMVServoController->ResetIMVStaticVariables() CMVServoController->Enable() SpontMode::Enter(fromModeId) ProcessAPRVPHighTimeTrigger() TriggerMgr::S_GetInstance()->EnableTrigger(eAPRVUpdateBRTrigger,APRV_UPDATE_BRTRIGGER) TriggerMgr::S_GetInstance()->EnableTrigger(eExtremeHigh_PHigh_Trigger) | Disable all trigger triggerMgrPtr->EnableIMVTrigger() triggerMgrPtr->DisableTrigger(eApneaAlarmTimeTrigger) Update phase setting Enable CMVServoController or VCVServoController base on current Ven Type Save log | PhaseMgr::S_GetInstance()->ProcessNewVentType(ePressureVentilation) CMVServoController->Enable() (PhaseMgr::S_GetInstance())->ProcessNewTriggerType(ePressureTrigger) PhaseMgr::S_GetInstance()->ProcessNewPhase(eExhPhase) TriggerMgr::S_GetInstance()->DisableTrigger(eBpTimeTrigger) UNSIGNED breathPeriod = TriggerMgr::S_GetInstance()->GetTimePeriod(eNIVBpTimeTrigger) TriggerMgr::S_GetInstance()->EnableTrigger(eNIVBpTimeTrigger,breathPeriod) (TriggerMgr::S_GetInstance())->DisableTrigger(ePtInhFlowTrigger) | PhaseMgr::S_GetInstance()->ProcessNewVentType(ePressureVentilation) check condition to : CMVServoController->ResetIMVStaticVariables() CMVServoController->Enable() PhaseMgr::S_GetInstance()->ProcessNewPhase(eExhPhase); | (PhaseMgr::S_GetInstance())->ProcessNewVentType(ePressureVentilation) Check condition to: CMVServoController->ResetIMVStaticVariables() CMVServoController->Enable() PhaseMgr::S_GetInstance()->ProcessNewPhase(eExhPhase) | PhaseMgr::S_GetInstance()->ProcessNewVentType(eVolumeVentilation) Check condition to : VCVServoController->ResetIMVStaticVariables() and ResetVolumeAdjustmentOfVCVInhPhase VCVServoController->Enable(); PhaseMgr::S_GetInstance()->ProcessNewPhase(eExhPhase) | PhaseMgr::S_GetInstance()->ProcessNewVentType(eVolumeVentilation) Check condition to : VCVServoController->ResetIMVStaticVariables() and ResetVolumeAdjustmentOfVCVInhPhase VCVServoController->Enable(); PhaseMgr::S_GetInstance()->ProcessNewPhase(eExhPhase) | PhaseMgr::S_GetInstance()->ProcessNewVentType(ePressureVentilation) if(fromModeId != eApneaVAMode) do something CMVServoController->Enable() PhaseMgr::S_GetInstance()->ProcessNewPhase(eExhPhase) Check Status = proxySensor->getLastStatus() to do something | InpedanceLine->SetState(eOn) InpedanceLine->SetDesiredState(eOff) SafetyValve->SetDesiredState(eClose) Disable all trigger EnableIMVTrigger CMVServoController->Enable() Check status of proxysensor to do something AlarmTask::S_GetInstance()->SetIsActive(eApnea, eActive) ProcessInhPhase() | PhaseMgr::S_GetInstance()->ProcessNewVentType(ePressureVentilation) check if(fromModeId != eApneaVAMode) to process somehthing CMVServoController->Enable() PhaseMgr::S_GetInstance()->ProcessNewPhase(eExhPhase) Check status of ProxySensor to process something | Enable the reset triggers Enable the total gas loss trigger Enable the alarm status update trigger ReportModeChangeInducedReset () Notify AlarmStat objects that AV is active | Put the ventilator into the eExhPhase phase Disable some triggers: eManualInhOpReqTrigger + ePtInhPressTrigger + ePtInhFlowTrigger Enable the Update Status trigger Enable the Reconected trigger Enable the reset trigger Update the alarm status to indicate that the total gas loss alarm is not active Update the Disconnected Circuit alarm status CMVServoController->Enable() Suspend auto-zero scheduling | SafetyValve->SetDesiredState(eOpen) triggerMgrPtr->DisableAllTriggers() triggerMgrPtr->DisableTrigger(eHardwareErrorAlarmTriger) triggerMgrPtr->EnableTrigger(eModeChangeTrigger) triggerMgrPtr->EnableTrigger(eHardwareErrorAutoresetTrigger) triggerMgrPtr->EnableTrigger(eAlarmResetOpReqTrigger) triggerMgrPtr->DisableTrigger(eUpdateStatusTimeTrigger) triggerMgrPtr->EnableTrigger(eUpdateStatusTimeTrigger, INIT_UPDATE_STAT_TIME) Suspend auto-zero scheduling HfoController->Disable() HfoServo::S_GetInstance()->Disable() PhaseMgr::S_GetInstance()->ProcessNewPhase(eAlarmPhase) | NE | NE | SafetyValve->SetDesiredState(eOpen) triggerMgrPtr->DisableAllTriggers() triggerMgrPtr->EnableTrigger(eModeChangeTrigger) triggerMgrPtr->EnableTrigger(eAlarmResetOpReqTrigger) triggerMgrPtr->DisableTrigger(eUpdateStatusTimeTrigger) triggerMgrPtr->EnableTrigger(eUpdateStatusTimeTrigger, INIT_UPDATE_STAT_TIME) CMVServoController->Disable() VCVServoController->Disable() HfoController->Disable() Suspend auto-zero scheduling (PdTask::S_GetInstance())->S_SetEventFlag(NONBREATHING_MODE_EVENT_FLAG) PhaseMgr::S_GetInstance()->ProcessNewPhase(eAlarmPhase)
Exit | (PdTask::S_GetInstance())->S_SetEventFlag(EXIT_STANDBYMODE_FLAG); SafetyValve->SetDesiredState(eOpen); PurgeFlow->SetDesiredState(eOn); PInhAzTestPtr->ResumeTest(); PExhAzTestPtr->ResumeTest(); HfoValve->EnableMotorErrorCheck(); O2Valve->EnableMotorErrorCheck(); AirValve->EnableMotorErrorCheck(); ExhValve->EnableMotorErrorCheck(); WatchDog::S_GetInstance()->EnableWatchDog(); | NA | Close the safety valvev Initialize pointer to the TriggerMgr Disable triggers Resume auto-zero scheduling | Disable triggers: eUpdateStatusTimeTrigger + eAlarmResetOpReqTrigger update CurrentTestId = eNoBdEventId Resume auto-zero scheduling | Resume auto-zero scheduling. | Disable the normal mode specific triggers Assume that HIP-2 will be checked. If going to a mode that does not check for the HIP-2 condition, it must override the trigger | NA | Disable eOcclAutoResetTrigger Update the Occlusion alarm status AlarmMode::Exit () | Disable the gas loss triggers Update the alarm status to indicate that the total gas loss alarm is not active Invoke the base class Exit() operation | NE | AirFlowController->Disable() O2FlowController->Disable() NCPAPController->Disable() SafetyValve->SetDesiredState(eOpen) | NA | NA | NA | NA | triggerMgrPtr->DisableTrigger(eAPRVLowPressTimeTrigger) triggerMgrPtr->DisableTrigger(eAPRVHighPressTimeTrigger) triggerMgrPtr->DisableTrigger(eExtremeHigh_PHigh_Trigger) triggerMgrPtr->DisableTrigger(eHPHighAlarmTrigger) CMVServoController->Disable() | Disable CMVServoController, VCVServoController Save Log | NA | NA | NA | NA | NA | TriggerMgr::S_GetInstance()->DisableTrigger(eProxyConnectTrigger) TriggerMgr::S_GetInstance()->DisableTrigger(eProxyDisconnectTrigger) SafetyValve->SetDesiredState(eOpen) CMVServoController->Disable() | AlarmTask::S_GetInstance()->SetIsActive(eApnea, eNotActive) MgrBreathData::S_GetInstance()->InitializeBreathDataPtr() TriggerMgr::S_GetInstance()->DisableTrigger(eBpBackupTimeTrigger) CMVServoController->Disable() | NA | Disable AV Mode specific triggers | Disable triggers: eReconnectionTrigger + eUpdateStatusTimeTrigger Update the Disconnected Circuit alarm status CMVServoController->Disable() Resume auto-zero scheduling | Update the High Press alarm status Disable triggers Resume auto-zero scheduling | NE | NE | Disable triggers: eHighMAP_2AutoResetTrigger + eUpdateStatusTimeTrigger Resume auto-zero scheduling
Go To another mode | None | NA | None | None | None | None | None | None | None | NE | None | None | None | None | None | Go to SponMode | None | Go to MixedMode | Go to MIxMode | Go to MixMode | Go to MixMode | Go to MixMode | Go to MixMode | None | go to MixMode | None | None | go to AlarmBreathingMode | NE | NE | go to AlarmBreathingMode

## List AlarmStat
Name | E_GUIEventID id | E_AlarmLevel initialPriority | E_AlarmLevel finalPriority | UNSIGNED tim | UNSIGNED occurent
eUsingDefaultSettings | eUsingDefaultSettings | eHigh | eHigh | 0 | 0
eSettingChange | eSettingChange | eLow | eLow | 0 | 0
eCorruptCompliance | eCorruptCompliance | eHigh | eHigh | 0 | 0
eCorruptAltitude | eCorruptAltitude | eHigh | eHigh | 0 | 0
ePrimaryAlarmFailure | ePrimaryAlarmFailure | eHigh | eHigh | 0 | 0
eO2MotorLostCommunication | eO2MotorLostCommunication | eHigh | eHigh | 0 | 0
eAirMotorLostCommunication | eAirMotorLostCommunication | eHigh | eHigh | 0 | 0
eExhMotorLostCommunication | eExhMotorLostCommunication | eHigh | eHigh | 0 | 0
eHFOMotorLostCommunication | eHFOMotorLostCommunication | eHigh | eHigh | 0 | 0
eO2MotorError | eO2MotorError | eHigh | eHigh | 0 | 0
eAirMotorError | eAirMotorError | eHigh | eHigh | 0 | 0
eExhMotorError | eExhMotorError | eHigh | eHigh | 0 | 0
eHFOMotorError | eHFOMotorError | eHigh | eHigh | 0 | 0
eAirValveStuckClosed | eAirValveStuckClosed | eHigh | eHigh | 0 | 0
eAirValveStuckOpen | eAirValveStuckOpen | eHigh | eHigh | 0 | 0
eO2ValveStuckClosed | eO2ValveStuckClosed | eHigh | eHigh | 0 | 0
eO2ValveStuckOpen | eO2ValveStuckOpen | eHigh | eHigh | 0 | 0
eExhValveStuckClosed | eExhValveStuckClosed | eHigh | eHigh | 0 | 0
eExhValveStuckOpen | eExhValveStuckOpen | eHigh | eHigh | 0 | 0
eDaemonLostSignal | eDaemonLostSignal | eHigh | eHigh | 0 | 0
eSpeakerError | eSpeakerError | eHigh | eHigh | 0 | 0
eAirFlowSensorError | eAirFlowSensorError | eMedium | eHigh | 0 | 2
eOxygenFlowsensorError | eOxygenFlowsensorError | eHigh | eHigh | 0 | 0
eBadBatteryCurrentSensor | eBadBatteryCurrentSensor | eMedium | eHigh | 0 | 0
eBadBatteryVoltageSensor | eBadBatteryVoltageSensor | eHigh | eHigh | 0 | 0
eBadInternalTempSensor | eBadInternalTempSensor | eHigh | eHigh | 0 | 0
eBadInternalO2Sensor | eBadInternalO2Sensor | eHigh | eHigh | 0 | 0
eBadADCWrapAISensor | eBadADCWrapAISensor | eHigh | eHigh | 0 | 0
eLowExternalBattery | eLowExternalBattery | eMedium | eMedium | 0 | 0
eLowInternalBattery | eLowInternalBattery | eHigh | eHigh | 0 | 0
ePreviousPowerFailed | ePreviousPowerFailed | eHigh | eHigh | 0 | 0
eWatchDogKickoutAlarm | eWatchDogKickoutAlarm | eHigh | eHigh | 0 | 0
eHip | eHip | eMedium | eHigh | 0 | 2
eExtremeHip | eExtremeHip | eHigh | eHigh | 0 | 0
eHighInspPressure | eHighInspPressure | eHigh | eHigh | ELEVATIONTIME | 0
eLowerLimitPressure | eLowerLimitPressure | eMedium | eHigh | 0 | 4
eLowInspPressure | eLowInspPressure | eMedium | eHigh | ELEVATIONTIME | 0
eOverLimitPressure | eOverLimitPressure | eMedium | eHigh | 0 | 4
eHighNCPAP | eHighNCPAP | eHigh | eHigh | 0 | 0
eLowNCPAP | eLowNCPAP | eHigh | eHigh | 0 | 0
eOcclusion | eOcclusion | eHigh | eHigh | 0 | 0
eGasSuppliesLost | eGasSuppliesLost | eHigh | eHigh | 0 | 0
eApnea | eApnea | eHigh | eHigh | 0 | 0
eLowPeep | eLowPeep | eMedium | eHigh | ELEVATIONTIME | 0
eHighPeep | eHighPeep | eMedium | eHigh | ELEVATIONTIME | 0
eExtremeHighPeep | eExtremeHighPeep | eHigh | eHigh | 0 | 0
eExtremeHigh_PHigh | eExtremeHigh_PHigh | eHigh | eHigh | 0 | 0
eAPRVHigh_PHigh | eAPRVHigh_PHigh | eMedium | eHigh | 0 | 2
eAPRVLow_PHigh | eAPRVLow_PHigh | eMedium | eHigh | ELEVATIONTIME | 0
eAPRVHigh_PLow | eAPRVHigh_PLow | eMedium | eHigh | ELEVATIONTIME | 0
eAPRVLow_PLow | eAPRVLow_PLow | eMedium | eHigh | ELEVATIONTIME | 0
eLowVte | eLowVte | eMedium | eHigh | ELEVATIONTIME | 0
eHighVte | eHighVte | eMedium | eHigh | ELEVATIONTIME | 0
eAirSourceFault | eAirSourceFault | eHigh | eHigh | 0 | 0
eLowO2Supply | eLowO2Supply | eHigh | eHigh | 0 | 0
eLowAirSupply | eLowAirSupply | eHigh | eHigh | 0 | 0
eUseO2Only | eUseO2Only | eHigh | eHigh | 0 | 0
eUseAirOnly | eUseAirOnly | eHigh | eHigh | 0 | 0
eHighO2 | eHighO2 | eHigh | eHigh | 0 | 0
eLowO2 | eLowO2 | eHigh | eHigh | 0 | 0
eHighRespiratoryRate | eHighRespiratoryRate | eMedium | eHigh | ELEVATIONTIME | 0
eLowMinuteVolume | eLowMinuteVolume | eMedium | eHigh | ELEVATIONTIME | 0
eHighMinuteVolume | eHighMinuteVolume | eMedium | eHigh | ELEVATIONTIME | 0
eLowVtSpontaneous | eLowVtSpontaneous | eMedium | eHigh | ELEVATIONTIME | 0
eLowVtMandatory | eLowVtMandatory | eMedium | eHigh | ELEVATIONTIME | 0
eITimeTooLong | eITimeTooLong | eMedium | eHigh | ELEVATIONTIME | 0
eHighEnclosureTemp | eHighEnclosureTemp | eHigh | eHigh | 0 | 0
eHighInternalO2 | eHighInternalO2 | eHigh | eHigh | 0 | 0
eBadFileSystem | eBadFileSystem | eHigh | eHigh | 0 | 0
eDisconnectedCircuit | eDisconnectedCircuit | eHigh | eHigh | 0 | 0
eOxygenCalibration | eOxygenCalibration | eLow | eLow | 0 | 0
e100percentO2Enable | e100percentO2Enable | eLow | eLow | 0 | 0
eFanError | eFanError | eHigh | eHigh | 0 | 0
eHFODoorOpen | eHFODoorOpen | eHigh | eHigh | 0 | 0
eHFOPistonNotConnect | eHFOPistonNotConnect | eHigh | eHigh | 0 | 0
eProxyDisconnect | eProxyDisconnect | eLow | eHigh | ELEVATIONTIME | 0
eDACLoopBackFailure | eDACLoopBackFailure | eHigh | eHigh | 0 | 0
eInhPAutoZero | eInhPAutoZero | eMedium | eHigh | 0 | 2
eExhPAutoZero | eExhPAutoZero | eMedium | eHigh | ELEVATIONTIME | 0
eFrontPanelDisconnect | eFrontPanelDisconnect | eMedium | eHigh | ELEVATIONTIME | 0
eProximalLeakage | eProximalLeakage | eMedium | eHigh | ELEVATIONTIME | 0
eHighMAP | eHighMAP | eHigh | eHigh | 0 | 0
eExtremeHighMAP | eExtremeHighMAP | eHigh | eHigh | 0 | 0
eAbnormalMAP | eAbnormalMAP | eHigh | eHigh | 0 | 0
eLowMAP | eLowMAP | eHigh | eHigh | 0 | 0
eExtremeLowMAP | eExtremeLowMAP | eHigh | eHigh | 0 | 0
eExtremeLowMAP | eExtremeLowMAP | eHigh | eHigh | 0 | 0
eHighAmplitude | eHighAmplitude | eHigh | eHigh | 0 | 0
eLowAmplitude | eLowAmplitude | eHigh | eHigh | 0 | 0
eProximalLeakage | eProximalLeakage | eMedium | eHigh | ELEVATIONTIME | 0
eETTubeDisconnect | eETTubeDisconnect | eHigh | eHigh | 0 | 0
eLowSIPress | eLowSIPress | eHigh | eHigh | 0 | 0
eHighSIPress | eHighSIPress | eHigh | eHigh | 0 | 0
eExtremeHighSIPress | eExtremeHighSIPress | eHigh | eHigh | 0 | 0
eCircuitLeakInHFO | eCircuitLeakInHFO | eHigh | eHigh | 0 | 0